## Slide 1
NanoVNA
Mark Braunstein WA4KFZ
Ole Virginia Hams
November 2024

## Slide 2: NanoVNA Overview
A low-cost vector network analyzer
  $60-$90 range Includes:
USB cable
SMA extension cables
SMA calibrations
  short
  open
  load
  thru
https://groups.io/g/nanovna-users
https://nanovna.com/

## Slide 3: NanoVNA Overview
Different models available
Frequency range starts at 50kHz
Some variants cover up to 3GHz
Measurements are made over 101 ‘points’
  Evenly distributed over the measurement’s Start and Stop frequency range
Dynamic range: 40-70dB (depending on frequency range)

## Slide 4: NanoVNA Display Modes
LOGMAG (log magnitude in dB)
Phase
Delay
Smith (Smith Chart)
SWR
Polar
Linear
Four traces can be displayed at the same time
Markers can be set for each trace

## Slide 5: NanoVNA Display Modes
Two-port Device Example
https://groups.io/g/nanovna-users
One-port SWR Example
https://groups.io/g/nanovna-users

## Slide 6: NanoVNA User Interface
Touchscreen operation for all other functions
https://groups.io/g/nanovna-users

## Slide 7: NanoVNA Typical Use Cases
https://groups.io/g/nanovna-users
Ant.
Port 1
Port 2
Port 1
Port 2
S11
“one-port”
measurement
S21
“two-port”
measurement
incident
reflected
incident
returned
Filter
7

## Slide 8: NanoVNA Use Case Warning
Several web sites, user forums, etc. indicate that the NanoVNA (and similar variants) do not have ESD (static) protection on the inputs
Directly connecting a large outdoor antenna can induce voltages large enough to damage the input
circuitry
Antenna protectors (Polyphaser, Alpha Delta, etc.) use gas discharge tubes (GDT) that do not offer the necessary low-voltage protection
ESD protection diode circuit
  https://www.qrpforum.de/index.php?attachment/23764-nanovna-input-protector-pdf/
NanoVNA protection video
  https://youtu.be/totwu4IbavE?si=gSg0mIKITo0l3UEb

## Slide 9: NanoVNA Test Port Cables
The use of short test port cables is HIGHLY recommended
Prevents accidental damage to the NanoVNA connectors
Reduces stress and wear on NanoVNA SMA connectors
https://groups.io/g/nanovna-users

## Slide 10: NanoVNA Display Software
Windows PC
https://groups.io/g/nanovnausers/wiki#Software
Android App

## Slide 11: NanoVNA Calibration
https://groups.io/g/nanovna-users
Calibration Steps | Measurements
Open | One & Two Port
Short | One & Two Port
Load | One & Two Port
Isolation | Two Port
Thru | Two Port
11
https://groups.io/g/nanovna-users

## Slide 12: NanoVNA Calibration
The calibration process compensates for frequency-dependent imperfections in the internal circuitry
https://groups.io/g/nanovna-users
12

## Slide 13: “I know you are but what am I…?”
There are multiple variants of the NanoVNA on the market
Some are “genuine” while others are poor quality clones
“Turf wars” for some of the variants regarding licensing

## Slide 14: Typical “Scary” Language…

## Slide 15: Conclusion
Overall, the NanoVNA is a good unit for general amateur radio use
  Best performance (based on dynamic range) is below 300MHz
Buy the NanoVNA from “approved” vendors
Be mindful of ESD precautions
Handle SMA connectors with care

## Slide 16: Backup – Technical Material

## Slide 17: What is a ‘VNA’?
VNA = Vector Network Analyzer
What’s a ‘vector’?
What’s a ‘network’?

## Slide 18: What’s a Vector?
A measure of the magnitude and angle of a signal in an AC circuit
In RF, the angle is a representation of the phase difference between two
signals
In RF, the magnitude is a representation of the amplitude difference between two signals
The phase and magnitude are measured at a single frequency

## Slide 19: What’s a Vector?
https://techweb.rohm.com/product/power-device/si/15939/
Resistance: current and voltage are always in phase
Reactance: current and voltage are 90 degrees out of phase

## Slide 20: What’s a Network?
A network is a “stimulus/response” representation of an electronic circuit
A network could consist of multiple components – we are analyzing the
cumulative behavior
Think of a network as a “black box” where you don’t know exactly what’s inside – you’re trying to figure out how it behaves by stimulating it with a signal and measuring the corresponding response

## Slide 21: What’s a Network?
https://en.wikipedia.org/wiki/Two-port_network
A generalized two-port network The voltage and current at each port can be independently measured
Typical network numbering convention:
Port 1 = input and Port 2 = output
Low frequencies
Types of low-frequency Network Parameters
H parameters (hybrid)
G parameters (inverse hybrid)
Z parameters (impedance)
Y parameters (admittance)
ABCD parameters (cascade)

## Slide 22: What’s a Network?
An S-parameter two-port network
The incident and reflected power at each port can be independently measured
Typical network numbering convention:
Port 1 = input and Port 2 = output
RF & Microwave frequencies
Notation
S21
Output at port 2 Due to input at port 1
S11
Output at port 1 Due to input at port 1
https://en.wikipedia.org/wiki/Two-port_network

## Slide 23: The ‘Bridge’ is the Key Component
https://electronicscoach.com/ac-bridges.html
General Case
The bridge is predominantly used in low frequency measurements
When the bridge is ‘balanced’ there is no difference signal at the Detector
and
occurs when:
At microwave frequencies, directional couplers
replace the ‘arms’ of the bridge circuit to
measure RF power ratios
23

## Slide 24: Bridge Concept Used in the NanoVNA
https://axotron.se/blog/vector-network-analyzer-a-christmas-present-to-myself/
The ‘fourth arm’ of the bridge connects to the TX port
The resistor in the middle supports the Detector
connection (mixer M1)
The local oscillator (LO) and stimulus (RF) oscillator are offset in frequency by a few kHz and measured at audio frequencies with soundcard-type components
Two-port thru measurements use the TX and RX ports to measure the one-way transmission parameters.
A one-port measurement at the TX port can still be performed, even if the device under test is connected between the TX and RX ports for a two-port measurement (i.e., you can measure both S11 and S21
with the same two-port configuration).
24